## Slide 1
CCD gain is expressed in units of electrons per count For example, a gain of 1.8 e-/count means that the camera produces 1 count for every 1.8 recorded electrons. It is important to note that every measurement you make in a CCD image uses units of counts. Since one camera may use a different gain than another camera, count units do not provide a straightforward comparison to be made. For example, suppose two cameras each record 24 electrons in a certain pixel. If the gain of the first camera is 2.0 and the gain of the second camera is 8.0, the same pixel would measure 12 counts in the image from the first camera and 3 counts in the image from the second camera. Without knowing the gain, comparing 12 counts against 3 counts is pretty meaningless.
Maths:
SC | The signal measured in count units in the CCD image
SE | The signal recorded in electron units by the CCD chip. This quantity is unknown.
NC | The total noise measured in count units in the CCD image.
NE | The total noise in terms of recorded electrons. This quantity is unknown.
g | The gain, in units of electrons per count. This will be calculated.
RE | The readout noise of the CCD, in [electrons]. This quantity is unknown.
s E | The photon noise in the signal NE
s o | An additional noise source in the image. This is described below.
https://www.mirametrics.com/tech_note_ccdgain.php
Signals and noise are related through the gain factor
SE=g*SC and NE=g*NC
Measured signal and noise:
SC =SE /g and NC =NE /g

## Slide 2
Noise analysis of CCD:
RE = Readout noise, σE = Photon (shot) noise, σ0,E =other noise
These noise source are independent and thus they add in quadrature, total noise
+
+
, for shot noise=
Let g be the gain, then total noise is
+
This is the equation of a line in which is the y axis, is the x axis, and the slope is 1/g. The extra terms are grouped together for the time being.
A better way to apply this equation is to plot measurements with as the y axis and as the x axis, as this gives the gain directly as the slope. Theoretically, at least, one could also calculate the readout noise,, from the point where the line hits the y axis at = 0.

## Slide 3
Crude Estimation of the Gain
Obtain images at different signal levels and subtract the bias from them. This is necessary because the bias level adds to the measured signal but does not contribute noise.
Measure the signal and noise in each image. The mean and standard deviation of a region of pixels give these quantities. Square the noise value to get a variance at each signal level.
For each image, plot Signal on the y axis against Variance on the x axis.
Find the slope of a line through the points. The gain equals the slope.
The mysterious extra noise term, , is attributable to pixel-to-pixel variations in the sensitivity of the CCD, known as the flat field effect. The flat field effect produces a pattern of apparently "random" scatter in a CCD image. Even an exposure with infinite signal to noise ratio ("S/N") shows the flat field pattern. Despite its appearance, the pattern is not actually random because it repeats from one image to another. (The link explain more details to measure this term)
https://www.mirametrics.com/tech_note_ccdgain.php
https://www.ptgrey.com/white-paper/id/10912
http://spiff.rit.edu/classes/phys445/lectures/gain/gain.html